Innomed Medical Inc.

Table of contents

1.	Checking requirements of technical data	3

1.1.	General requirements	3

1.2.	Instruments and devices recommended for checking the technical data	3

2.	MEASUREMENT PROCEDURES	4

2.1.	Safety technology tests	4

2.1.1.	Electrical strength test	4

2.1.2.	Earth leakage current measurement	4

2.1.3.	Patient leaking current (ECG)	5

2.1.4.	Patient leakage current generated by external voltage (ECG)	6

2.2.	Test of the amplifier	6

2.2.1.	Lead selector test with ECG simulator	6

2.2.2.	Lead selector measurement with sine signal	7

2.2.3.	Measurement of the input noise and check of the zero line	8

2.2.4.	Check of the CMRR	8

2.3.	Check of the PC connection	9

2.4.	Check of the clock settings and the memory contents	9

2.5.	Calibration of the touchscreen (if have a touchscreen)	9

# Checking requirements of technical data

## General requirements

During the measurements, a through care must be taken in order to properly form the measuring premise mechanically and electrically. The measuring room must meet the requirements for the operation of the device.

The clear understanding and measurement of the technical data are summarized in this chapter.

| WARNING! |  |
| --- | --- |
| For measurements or tests, the device’s own accessories must be used! |

| WARNING! |  |
| --- | --- |
| During the tests and/or measurements, the safety regulations must be strictly observed! |

| WARNING! |  |
| --- | --- |
| Unless other instruction, the tests must be done without baseline filter. |

## Instruments and devices recommended for checking the technical data

| 1 pc | High-voltage Breakdown SMU test device |  |
| --- | --- | --- |
| 1pc | Safety Tester (Metron QA-90, Fluke ESA620 Safety Tester or equivalent ) | Calibrated |
| 1 pc | Storage oscilloscope (2 mV/div, DC-20 MHz, Rbe  1 Mohm) |  |
| 1 pc | V.L.F. generator ( sine, tetragon, 0.1 mV pp - 5 Vpp, 0.01 - 500 Hz, | Calibrated |
| 1 pc | Multi-meter (0.05 Hz - 1 kHz, 1 mV, 1 %, Peak, RMS) |  |
| 1 pc | ECG simulator , Phantom 320 | Calibrated |
| 1 pc | PC with Innobase program |  |
| 1 pc | Measuring device for measuring leakage current |  |

Measuring device for measuring leakage current (MD)

# MEASUREMENT PROCEDURES

| WARNING! |  |
| --- | --- |
| Concerning the hereunder unregulated measurement procedures, the IEC601-1 Safety standard and the IEC 601-2-25 functional standard for ECG devices are normative. |

## Safety technology tests

| WARNING! |  |
| --- | --- |
| In case of repeated testing, the high voltage test shall be performed with reduced voltage. The below given voltages ares 80% of the nominal values required by the standard. |

### Electrical strength test

When performing the test, the voltage is gradually increased from 0 to the specified test voltage. The device is suitable if it can withstand the high voltage load for 20 seconds.

| Between the connected ECG cable | and | the mains cable with connected neutral and phase (except protective earth) | 3200Veff |  |
| --- | --- | --- | --- | --- |
| Between the he connected ECG cable | and | the connected USB cable | 3200Veff |  |
| Between the protective earth | and | connected neutral and phase | 1200Veff |  |
| Between the connected USB cable | and | the mains cable with connected neutral and phase (except protective earth) | 1200Veff |  |

Earth leakage current measurement

### Earth leakage current measurement

Earth Leakage Current: Normal Condition

•	Applicable to Class 1, Types B, BF and CF.

•	Measures earth leakage current of the instrument under test connected to the mains power supply; normal and reversed polarity using S2.

•	For Type BF and CF equipment, measures with the AP/GND switch S3 open and closed.

•	Maximum leakage current: 500 µA (Range: DC and AC up to 1 kHz).

Earth Leakage Current: S.F.C. Open Supply

•	Applicable to Class 1, Types B, BF and CF.

•	Measures earth leakage current of the instrument under test, with one open supply lead interrupted (S1 = open); normal and reversed polarity using S2.

•	For Type BF and CF equipment, measures with the AP/GND switch S3 open and closed.

•	Maximum leakage current: 1000 µA (Range: DC and AC up to 1 kHz).

|  | Limit | Measured value |
| --- | --- | --- |
| Without error condition | 500uA |  |
| One error condition (open supply) | 1000uA |  |
| One error condition (R.M. open supply) | 1000uA |  |

### Patient leaking current (ECG)

Patient Leakage Current: Normal Condition

•	Applicable to Classes 1 and 2, Types B, BF and CF.

•	Measures patient leakage current to earth from all Applied Parts in parallel; normal and reversed polarity using S2.

•	Maximum leakage current Types B and BF: 100 µA (Range: DC and AC up to 1 kHz). For  Type CF: 10 µA (Range: DC and AC up to 1 kHz).

Patient Leakage Current: S.F.C. Open Supply

•	Applicable to Classes 1 and 2, Types B, BF and CF.

•	Measures patient leakage current to earth from all Applied Parts in parallel with one supply lead interrupted (S1 = open); normal and reversed polarity using S2.

•	Maximum leakage current Types B and BF: 500 µA (Range: DC and AC up to 1 kHz). For Type CF: 50 µA (Range: DC and AC up to 1 kHz).

Patient Leakage Current: S.F.C. Open Earth

•	Applicable to Class 1, Types B, BF and CF.

•	Measures leakage current to earth from all Applied Parts in parallel with the Protective Earth open circuit (S4 = open); normal and reversed polarity using S2.

•	Maximum leakage current Types B and BF: 500 µA (Range: DC and AC up to         1 kHz). For Type CF: 50 µA (Range: DC and AC up to 1 kHz).

|  | Limit | Between protective earth and ECG |
| --- | --- | --- |
| Without error condition | max. 10uA |  |
| One error condition (open supply) | max. 50uA |  |
| One error condition (R.M. open supply) | max. 50uA |  |
| One error condition (open Earth) | max. 50uA |  |

### Patient leakage current generated by external voltage (ECG)

Patient Leakage Current: S.F.C. Mains on Applied Part

•	Applicable to Classes 1 and 2, Types BF and CF.

•	Measures leakage current to earth from Applied Part to earth caused by external mains volt-age on Applied Part, and with switch S5 open and closed.

•	Each polarity combination possible is tested using S5 and S6.

•	Maximum leakage current Type BF: 5000 µA (Range: DC and AC up to 1 kHz).  For Type CF: 50 µA (Range: DC and AC up to 1 kHz) (100 µA for CF defib paddles: IEC 601-2-4).

|  | Limit | Patient leakage current |
| --- | --- | --- |
| One error condition (open supply) | max. 50uA |  |
| One error condition (R.M. open supply) | max. 50uA |  |
| One error condition (open Earth) | max. 50uA |  |

## Test of the amplifier

It is sufficient to perform one of the 2.2.1 and the 2.2.2 measurements

### Lead selector test with ECG simulator

Polarity and amplitude test.

The patient cable of the device must be connected to the appropriate output of the ECG simulator (eg Phantom 320). The generator is in the 60 / min normal sine rhythm position. Sensitivity control * 1. The direction of the diversion must be checked on the record.

Suitable if the direction and size of the recorded signal are the same as specified in the MEASURING SHEETS section. The permissible error is +/- 3%.

|  | Lead selector measurement the 10 leads connected to the Phantom320 |
| --- | --- |
| Lead | Nominal amplitude of an ECG wave [mV] | Measured amplitude [mV] |
| I | +14,1 |  |
| II | +23,9 |  |
| III | +10,4 |  |
| aVR | -18,9 |  |
| aVL | +3,5 |  |
| aVF | +17,0 |  |
| V1 | -22,0 |  |
| V2 | -5,6 |  |
| V3 | +10,7 |  |
| V4 | +21,4 |  |
| V5 | +17,8 |  |
| V6 | +14,3 |  |

| WARNING! |  |
| --- | --- |
| When measuring with the Phantom 320 simulator, the amplitudes of the generator must first be determined by reference measurement. For series measurements, it must be compared to the recorded reference. |

### Lead selector measurement with sine signal

Use the patient cable of the device connected. A 10 Hz sine signal with 1 mVpp amplitude must be  connected between the single tested lead and all other leads connected to each other.

We measure the amplitudes on the printed record.

The permissible error is +/- 3%.

| Tested lead |  | I. | II. | III. | V1 | V2 | V3 | V4 | V5 | V6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R | Nominal | 1mV | 1mV | 0mV | 333uV | 333uV | 333uV | 333uV | 333uV | 333uV |
|  | Measured |  |  |  |  |  |  |  |  |  |
| L |  | 0mV | 0mV | 1mV | 333uV | 333uV | 333uV | 333uV | 333uV | 333uV |
| F |  | 1mV | 0mV | 1mV | 333uV | 333uV | 333uV | 333uV | 333uV | 333uV |
| C1 |  | 0mV | 0mV | 0mV | 1mV | 0mV | 0mV | 0mV | 0mV | 0mV |
| C2 |  | 0mV | 0mV | 0mV | 0mV | 1mV | 0mV | 0mV | 0mV | 0mV |
| C3 |  | 0mV | 0mV | 0mV | 0mV | 0mV | 1mV | 0mV | 0mV | 0mV |
| C4 |  | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 1mV | 0mV | 0mV |
| C5 |  | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 1mV | 0mV |
| C6 |  | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 0mV | 1mV |

### Measurement of the input noise and check of the zero line

The tested patient cable of the device is connected to the appropriate output of the ECG simulator (eg. Phantom 320). The simulator must be switched off.

In AUTO mode, a cca. 10 sec length record must be created with sensitivity control * 2, 25mm/sec speed for all group. The filters must be off.

The input noise is adequate if the peak-to-peak noise level does not exceed 0.4 mm at any lead.

A zero line is adequate if the deviation from the baseline is less than 4 mm anywhere on the register and display.

|  | Input noise and zero line offset measurement the lead 10 is connected to the Phantom320, the simulator is turned off. |
| --- | --- |
| Lead | Maximum value of noise [mm] | Measured value [mm] | Zero line offset maximum [mm] | Measured value [mm] |
| I | 0,4mm |  | 4mm |  |
| II | 0,4mm |  | 4mm |  |
| III | 0,4mm |  | 4mm |  |
| aVR | 0,4mm |  | 4mm |  |
| aVL | 0,4mm |  | 4mm |  |
| aVF | 0,4mm |  | 4mm |  |
| V1 | 0,4mm |  | 4mm |  |
| V2 | 0,4mm |  | 4mm |  |
| V3 | 0,4mm |  | 4mm |  |
| V4 | 0,4mm |  | 4mm |  |
| V5 | 0,4mm |  | 4mm |  |
| V6 | 0,4mm |  | 4mm |  |

### Check of the CMRR

Place the ECG device on a metal plate connected to the protective earth. Connect 20Vpp
 	60Hz signal from the sinusoidal generator between al ECG leads and the metal plate.

The ECG device must be connected to the mains.

In AUTO mode, a cca. 10 sec length record must be created with the sensitivity control * 1,  	25mm/sec speed for all group. All filters must be off.

Suitable if the signal is not greater than 2 mm (100dB)

|  | CMRR measurement |
| --- | --- |
| Lead | maximum amplitude of the wave [mm] | Measured amplitude [mm] |
| I | 2 mm |  |
| II | 2 mm |  |
| III | 2 mm |  |
| aVR | 2 mm |  |
| aVL | 2 mm |  |
| aVF | 2 mm |  |
| V1 | 2 mm |  |
| V2 | 2 mm |  |
| V3 | 2 mm |  |
| V4 | 2 mm |  |
| V5 | 2 mm |  |
| V6 | 2 mm |  |

## Check of the PC connection

Connect the ECG simulator to the input of the device. On the PC, monitoring the signal and switch through all gain stages, all speed stages, the filter and all leads. Check that the setting and signal are seen on the PC match the settings are seen on the device. Then start automatic storage on the PC, store it and retrieve it from the database to check the quality of the signal on the PC display. Check the R-wave amplitudes on each leads on both the device and PC display.

|  | PC connection is good |
| --- | --- |
| Check PC connection via USB |  |
| Check PC connection via Bluetooth ( if available) |  |

Print the signal from the PC to a laser printer and check the quality of the ECG signal. Check the parameters of the analysis, compare the numerical data seen on the display with the amplitude and time parameters that can be read from the recorded curves  printed with the device and as appear on the printout of the laser printer.

If the device has optional BT connection, it should also be checked as specified later.

|  | Printing to a laser printer |
| --- | --- |
| Check PC connection via USB |  |
| Check PC connection via Bluetooth ( if available ) |  |

## Check of the clock settings and the memory contents

Check the clock setting

Check the storage

## Calibration of the touchscreen (if have a touchscreen)

Follow the instructions for use of the device to calibrate the touchscreen if necessary.

Documentation of responsibility and changed parts:

| Device type: |  |
| --- | --- |
| Device serial number: |  |
| Test was carried out by: |  |
| Date of the test: |  |

| Removed components: | Serial numbers: |
| --- | --- |
| Built in new components: | Serial numbers |

Signature:

............................................................…..

Name of person who made the test

# Text boxes

ECG Devices

Measuring Test  Report

Voltage meter b)